IN-LINE ELECTRICAL METROLOGY FOR HIGH-K GATE DIELECTRICS DEPOSITED BY ATOMIC LAYER CHEMICAL VAPOR DEPOSITION

H. De Witte1,(, S. Passefort2a, W. Besling3a, J.W.H. Maes1, K. Eason2b, E. Young4,

Z.M. Rittersma3b, M. Heyns5

1ASM International, Jan van Eycklaan 10, 3723 BC Bilthoven, The Netherlands

2aKLA-Tencor Corporation, 14 rue du Bois Sauvage, F-91000 Evry, France

2bKLA-Tencor Corporation, 160 Rio Robles, San Jose, CA 95134, USA

3aPhilips Research, Prof. Holstlaan, Eindhoven, The Netherlands

3bPhilips Research Leuven, Kapeldreef 75, 3001 Leuven, Belgium

4International Sematech, 2706 Montopolis Drive, Austin, Texas 78741, USA

5IMEC, Kapeldreef 75, B-3001 Leuven, Belgium

The use of the KLA Tencor Quantox in-line electrical metrology tool for evaluation and characterization of Atomic Layer Chemical Vapor Deposition ALCVDTM high k dielectrics has been investigated. Equivalent Oxide Thickness obtained from Quantox agrees very well with High Frequency Capacitance - Voltage, C-V, results. Also the Quantox ACTIVTM Index parameter shows nice correlation with CV Jgate. Therefore Quantox can be used as fast and much less expensive evaluation tool for optimization of ALCVDTM process conditions and process stability monitoring.

# INTRODUCTION

Industry roadmaps for CMOS technology indicate that conventional thermal silicon oxide gate dielectrics rapidly will run out of steam because leakage current requirements cannot be met. This issue can be addressed by introducing high-k dielectrics such as Al2O3, HfO2, ZrO2 or possibly their aluminates or silicates. Fully functional transistors with high-k films deposited by Atomic Layer Chemical Vapor Deposition (ALCVDTM []) are already reported by various groups [-,,,].

Introduction of high-k gate stacks also means appearance of new film metrology challenges. Accurate film thickness measurements and rapid turn around time electrical analysis are highly desirable for high-k process development, tuning and monitoring. In this paper we will evaluate characterization of ALCVD™ high-k gate dielectrics using a KLA-Tencor Quantox in-line electrical metrology tool. Capacitance, leakage current and interface state density measurements will be demonstrated for ALCVD™ films with equivalent oxide thicknesses down to 12Å.

# EXPERIMENTAL PROCEDURES

Typical ZrAlOx and HfO2 based high-k gate dielectrics were made with an ASM Polygon™ 8200 [1] gate stack cluster tool. To compare Quantox measurements with MOSCAP C-V a set of wafers was made with ZrAlOx high-k stacks. A 5Å SiO2 interface layer was grown on p-type device quality wafers, followed by ALCVD™ deposition of the high-k layer, annealing in N2 and in forming gas. Three different ZrAlOx layer thickness (30, 50, 100 ALCVD™ cycles) were investigated. Duplicate wafers with the same high-k stacks were made and further processed into capacitors with TiN electrodes. For the application part of this paper, the exact wafer description is included further on.

In-line electrical measurements were performed using the KLA-Tencor Quantox. The Corona-Oxide-Silicon (COS) Quantox system [, ] is based on combining three non-contacting technologies: charged corona, vibrating Kelvin probe and a pulsed light source, as shown on the left hand side of Fig. 1. Charged corona ions provide biasing and emulate the functions of the Metal Oxide Semiconductor (MOS) electrical contact. The pulsed light source enables the stimulus and detection of surface photo-voltage (SPV) and provides measurement of silicon band-bending and a direct measurement of Vfb. The result of combining these technologies is a system capable of generating Q-V and SPV-V curves that are analogous to traditional MOS C-V measurements as presented on the right hand side of Fig. 1. The flatband is determined using the Q-V and the Q-SPV data together [Vfb = Vsurf @ SPV = 0]. The Equivalent physical Oxide Thickness (EOT) parameter is determined from dielectric capacitance. The capacitance is determined from dQ/dV in accumulation in the COS system. The capacitance is converted to thickness using r = 3.9. In actual application, some second order corrections are applied to account for semiconductor band-bending, in partial accumulation.

Fig. 1. Left: Quantox COS measurement theory; Right: Q-V-SPV Sweep Parameter Extraction: The surface voltage is plotted on the left-hand ordinate and surface photo voltage is plotted on the right-hand ordinate.

Capacitance - Voltage (C-V) measurements were performed on capacitors of 40 m x 40 m with TiN electrode. The C-V EOT was extracted using the CVC modeling program from NCSU [], accounting for, band bending, depletion of gate electrode and quantum mechanical (QM) effects in substrate (optional).

# RESULTS AND DISCUSSION

First of all, to validate Quantox measurements on ALCVD™ high k layers, we start with a comparison with high frequency CV data on duplicate wafers focussing on the most important parameters for high k characterization such as capacitance, EOT extraction and leakage. Secondly, we will demonstrate a selection of three fields of application were the Quantox can be used for quick electrical characterization. The chosen applications consist of the determination of k values, the effect of post anneals as an example of process screening, and ALCVD™ process monitoring.

In Fig. 2 the Quantox C-V curves are compared with the high frequency C-V curves as measured on the TiN metal electrode capacitors. Very similar C-V curves are obtained with both techniques. Our major interest here concerns the accumulation capacitance, which is the raw data to extract the EOT. One clearly observes that the Quantox accumulation capacitance increases with decreasing ZrAlOx thickness and shows very similar values as the high frequency C-V data. Surprisingly, despite the quasistatic nature of the Quantox measurement, charge leakage does not seem to disturb C-V measurement on these layers. Further, one can clearly see a shift of approximately 1V on the voltage axis. This can mainly be explained by the different work function (Quantox Kelvin Probe is calibrated on pure oxides ~ -0.2 eV versus TiN electrodes for the C-V measurements).

Fig. 2. Quantox C-V curves (left) versus MOSCAP high frequency C-V curves (right) of ZrAlOx based high-k stacks of three different thicknesses.

The standard method in high frequency C-V to extract the EOT values from the measured C-V curves is by means of the Hauser fitting program [8]. This program can generate two kinds of numbers, depending on whether or not a QM correction is applied to the data. The Quantox C-V (C=dQ/dVsurf versus Vsurf) data can be modelled well by the Hauser method. Table 1 lists the obtained EOT (with QM correction) and COT (accumulation capacitance without QM correction) and compares them with the Quantox “GateToxTM” [] results. The latter are obtained with a high-speed Quantox measurement that uses a reduced set of data, combined with appropriate analysis algorithms, to extract the electrical thickness. It has been developed especially for ultra-thin advance gate process control (sub 30Å). Note that the GateTox™ values do not include a correction for the Quantum mechanical effect in the substrate.

Table 1 clearly shows that the GateTox™ data are in good agreement with the COT results obtained with the Hauser fitting. The Quantox Hauser EOT results are also compared to the EOT values obtained with the MOSCAPs. Also here excellent agreement is obtained for all three layers. The 30 cycles ZrAlOx film shows an EOT of 12.1 Å, demonstrating that ALCVD™ films with the thicknesses required for CMOS gate stack applications can be characterized very well using in-line electrical metrology. Note that C-V measurements on 15-20 Å SiO2 based gate oxides already are seriously hindered by charge leakage during the measurement. The result in this paper illustrates the leakage current improvement achieved by using oxides such as ZrAlOx and HfO2 (see application part of this paper).

|  | ASET-F5 Ellipsometer | MOSCAP | Quantox |
| --- | --- | --- | --- |
| ALCVD™ cycles | Optical Thickness (Å) | RI (633 nm) | EOT Hauser (Å) | EOT Hauser (Å) | COT Hauser (Å) | GateTox™ (Å) |
| 30 | 35.9 | 1.82 | 12.1 | 12.2 | 16.2 | 16.4 |
| 50 | 56.5 | 1.83 | 18.2 | 18.3 | 22.3 | 22.1 |
| 100 | 109 | 1.86 | 30.3 | 31.8 | 35.8 | 36.1 |

Table 1. Ellipsometry, high frequency C-V data and Quantox results of ZrAlOx based high-k gate stacks.

Besides the need for a quick method to measure the EOT, also a fast method to determine the physical (optical) thickness is required. For this purpose, KLA developed a special recipe on the ASET-F5 ellipsometer, which allows measuring thickness and Refractive Index (RI), giving an indication of the composition of the ZrAlOx layers. The results of these measurements are also listed in Table 1. Combination of EOT and optical thickness allows the determination of k values, which will be illustrated further on in this paper.

Gate leakage has been sited as one of the primary drivers for industry transition away from SiO2 as gate dielectric and toward high k as a replacement candidate. Thus, it is very important to attempt to probe the film leakage characteristics using the Quantox tool. The Quantox offers the “ACTIV™ Index” parameter to monitor this critical dielectric property. The parameter implements KT proprietary measurement and analysis algorithms to provide a quantification of the oxide integrity and quality. ACTIV™ Index correlates well with film leakage measured by MOSCAP and MOS transistor.

In Fig. 3 the Quantox ACTIV™ Index is plotted against the logarithm of the standard leakage parameter Jgate from MOSCAPs. A linear correlation is obtained, thereby confirming the Quantox ACTIV™ Index parameter as a leakage current monitor.

Fig. 3. Quantox ACTIVTM Index versus Jgate from High Frequency CV.

After this validation of the most important Quantox parameters for high k characterization we will now present some applications for ALCVD™ high k process development. First, a demonstration of k value extraction; secondly, the effect of post anneal treatments on a ALCVD™ high k layer and third, the use of the Quantox as in-line metrology for process stability monitoring.

The k values are determined by combining either Quantox or MOSCAP EOT values and optical thickness (ASET-F5) according to the following formula:

(1)

In Fig. 4 this method is illustrated for a thickness series of ALCVD™ ZrAlOx samples (25, 50, 100 and 200 cycles) with a Zr:Al ratio of 4:1 and a bottom layer of 15 cycles Al2O3. The slope of this curve gives a value of 20.8 as the k value for the ZrAlOx layer.

Fig. 4. Determination of ZrAlOx k value as the slope of this graph for a series of ZrAlOx layers with a Zr:Al ratio of 4:1.

In Fig. 5 the k values are given as a function of the Al2O3 composition for a series of thick ALCVD™ ZrAlOx layers (~ 1000 cycles total) with different Zr:Al ratio (1:2, 1:1, 4:1) and a 15 cycles Al2O3 bottom layer. In this case the k value of the whole stack is presented. The composition is determined by physical characterization, while the k values are determined either from Quantox or high frequency C-V combined with ASET-F5 optical thickness data. One observes a nearly linear scaling of the k values with the composition. As expected from the above validation experiment good agreement is obtained between Quantox and C-V data. Combination of this k-value data with crystallization temperature data allows quick selection of the material with the optimal composition for gate stack or other applications.

Fig. 5. Quantox and MOSCAP determined k values in function of the Al2O3 composition (physical characterization) for a series of ALCVD™ ZrAlOx films.

The second application is an example of using the Quantox for process development purposes. An as-deposited HfO2 layer (85 cycles ALCVD™ HfO2) on chemical oxide [] is compared to an equivalent layer with oxygen post anneal (both with forming gas anneal). The data shown in Table 2 indicate that the interfacial oxide growth during the anneal is rather limited. The flatband voltage Vfb is shifted to lower values, suggesting a decrease in the negative charges. Also the layer quality is improved by the anneal, as can be derived from the Dit, indicating a better interface quality, and a higher ACTIV™ Index, representative for lower leakage.

|  | GateToxTM [Å] | Vfb [V] | Dit x E10 [#/(eV_cm2)] | ACTIVTM Index |
| --- | --- | --- | --- | --- |
| As deposited | 17.9 | 1.225 | 59 | 3.769 |
| After anneal | 19.8 | 0.977 | 33 | 4.138 |

Table 2. Comparison of important Quantox parameters for the HfO2 layer before and after oxygen anneal.

As the Quantox is already extensively used as in-line metrology for monitoring process stability of SiO2 or SiOxNy based gate oxides, we now also started using it for ALCVD™ process monitoring. We therefore chose a deposited 100 cycles HfO2 film on chemical oxide, with no post anneals. The following parameters are monitored: the electrical thickness (GateTox™), the interface quality by means of the interface trap density Dit and flatband voltage Vfb, and the leakage by the ACTIV™ Index. The obtained results are graphically represented in Fig. 6 and the 1 sigma (%) over time for the different parameters under monitoring are shown in Table 3.

Fig. 6. ALCVD™ process stability monitoring by Quantox.

|  | GateToxTM | Activ™ Index | Dit | Vfb |
| --- | --- | --- | --- | --- |
| 1 sigma (%) | 2.8 | 2.4 | 14.6 | 7.2 |

Table 3. 1 sigma (%) variation over time for the Quantox parameters shown in Fig. 6.

# CONCLUSIONS

The present paper demonstrated the capabilities of the in-line electrical metrology tool for the characterization and process optimization and selection of ALCVD™ high k stacks. Capacitance and leakage current measurements have been demonstrated for high-k films with an electrical thickness as low as 12 Å.

Quantox measurement data are found to correlate well to MOSCAP results on the following parameters: first, Quantox GateTox™ to high frequency COT and secondly, Quantox ACTIV™ Index to Jgate (measured at Vfb + 1V) from the high frequency C-V measurements.

Quantox has been shown to be very useful for process applications, such as fast determination of k values, effects of oxygen post anneals as an example of process screening and process monitoring.

# REFERENCES

# Footnotes

( E-mail: hilde.de.witte@asm.com	Phone: +32-16-288244	Fax: +32-16-281221

# Endnotes

. ALCVDTM, PolygonTM 8200 are trademarks of ASM International nv.

. D.A. Buchanan, E.P. Gusev, E. Cartier, H. Okorn-Schmidt, K. Rim, M.A. Mocuta, A. Ajmera, M. Copel, S. Guha, N. Bojarczuk, A. Callegari, C.D. Emic, P. Kozlowski, K. Chan, R.J. Fleming, P.C. Jamison, J. Brown, R. Arndt, Electron Devices Meeting 2000, IEDM Technical Digest. International, (2000) 223.

. K. Torii, Y. Shimamoto, S. Saito, K. Obata, T. Yamauchi, D. Hisamoto, N. Yokoyama, M. Hiratani, and T. Onai, International Workshop on Gate Insulator 2001, (2001) 8-06.

. Y.H. Kim, C.H. Lee, T.S. Jeon, W.P. Bai, C.H. Choi, S.J. Lee, L. Xinjian, R. Clarks, D. Roberts, D.L. Kwong, D.L., Electron Devices Meeting 2001, IEDM Technical Digest. International, (2001) 30.5.1.

. J.M. Hergenrother, G.D. Wilk, T. Nigam, F.P. Klemens, D. Monroe, P.J. Silverman, T.W. Sorsch, B. Busch, M.L. Green, M.R. Baker, M.R. Baker, T. Boone, M.K. Bude, N.A. Ciampa, E.J. Ferry, A.T. Fiory, S.J. Hillenius, D.C. Jacobson, Johnson, Electron Devices Meeting, 2001, IEDM Technical Digest. International, (2001) 3.1.1.

. T. G. Miller, Semiconductor International (1995).

. J. Guan et al., ECS Meeting, Abstracts, MA 99-2, (1999) 1106.

. J.R. Hauser and K. Ahmed, Characterization and Metrology for ULSI Tech.: Int. Conf., AIP Conference Proceedings, 449 (1), (1998) 235.

. GateToxTM and ACTIV™ Index are trademarks of KLA- Tencor corp.

. M.M. Heyns, T. Bearda, I. Cornelissen, S. De Gendt, R. Degraeve, G. Groeseneken, C. Kenens, D.M. Knotter, L.M. Loewenstein, P.W. Mertens, S. Mertens, M. Meuris, T. Nigam, M. Schaekers, I. Teerlinck, W. Vandervorst, R. Vos, K. Wolke, IBM Journal of Research and Development, 43 (3), (1999) 339.